Audit forms and check sheets for stack emissions monitoring

It is strongly recommended that regular audits of MCERTS accredited stack emissions monitoring organisations are carried out to ensure the accredited organisations continue to meet the requirements of MCERTS.

We strongly recommend that where possible audits are carried out on an unannounced basis. This gives the truest reflection of how work is normally carried out.

The audit forms are in Word format, so they can be adapted by the auditor. To help the auditor, notes in italics have been added to the forms.

The forms are:

A1: Operator, auditor, monitoring organisation and MCERTS documentation

A2: Health and safety

A3: Isokinetic sampling, wet chemistry and multiphase sampling

A4: Instrumental methods

Check sheets have also been included for checking site specific protocols and monitoring reports.

If you have any comments on the audit forms or check sheets please contact Rupert Standring at rupert.standring@environment-agency.gov.uk

| A1: Operator, auditor, monitoring organisation and MCERTS documentation |  |
| --- | --- |
| Operator |  |
| Release point(s) | Unique stack identification from Permit. |
| Auditor(s) |  |
| Monitoring organisation |  |
| Accreditation | Check that monitoring organisation has MCERTS accreditation for the sampling method and relevant supplementary parameters, such as oxygen. |
| Team leader | At least one member of the team must have MCERTS Level 2 plus the Technical Endorsements relevant to the methods required. TE 1 particulates; TE 2 multi phase (e.g. dioxins); TE 3 gases by manual methods; TE 4 gases by instrumental methods. |
| Team member | Must be a Level 1, 2 or registered as a trainee. |
| Site review | Date it was carried out. Was it done by a site visit? |
| Site specific protocol (SSP) | Does it meet MCERTS requirements? (see template in MCERTS performance standard). Does it match the permit? (e.g. methods listed, reference conditions). |
| Deviations from SSP | Are there any requirements in the SSP that the test team believe they will not be able to follow? If so, ask them to explain and justify deviations. |
| Methods manual | Check the monitoring organisation’s methods manual is available on site. |

| A2: Health and safety |  |
| --- | --- |
| Permit to work |  |
| Risk assessment | Must be signed off by each team member on each day. |
| COSHH assessment | Should be available for chemicals brought to site by monitoring organisation and for chemicals in the stack gas. |
| PPE requirements | Have the monitoring team followed the site rules? Safety spectacles required while leak checking glassware and handling chemicals (e.g. solvents and solutions). Appropriate chemically resistant gloves for handling chemicals. Appropriate gloves for handling equipment and climbing ladders. |
| First aid kit | Check that an in date eye-wash solution is available. |
| Platform details | Have the monitoring team checked if there is a permanent platform inspection certificate available before going up a platform? Have the monitoring team checked that temporary scaffold has been inspected and approved? (i.e. Scaftag system). For most monitoring work scaffold should be heavy duty and fixed to a permanent structure. See TGN M1 for further information. |
| Emergency rescue plan | Have the monitoring team checked if there is an emergency rescue plan in place? |
| Work area | Is it tidy? Check for trip hazards from trailing leads. |
| Exclusion zone | Work area below a platform must be clearly cordoned off. Men working overhead signs in place. |
| Lifting system | Observe how the equipment is lifted to the working area? Is it done manually? Is a portable lifting system used? Is it secured to a suitable structure? Is the equipment showing signs of wear and tear? Have personnel been on lifting and slinging courses? |
| Port sealing / gas venting | Check how the sampling access port is sealed during sampling? If working in an enclosed area, is the sample gas vented safely? |
| Communications | The team may need to communicate when working in different areas (i.e. ground level and platform). How do they communicate? Telephone, radio, shouting. Are man down alarms used? Has communication been established with staff responsible for operating the process? (e.g. control room). |
| Working hours & rest periods | Is the work plan sensible? Has it taken driving to and from site into account? Is artificial lighting available for work when it’s dark? |
| Chemical storage | Do the monitoring team have fridges, solvent cabinets and cool boxes? Are items clearly labelled? Check use by dates on solutions and solvents. |
| Cylinder storage | Are all the cylinders secured? (i.e. tied up or held in racks). Check cylinders are not transported with regulators on. Check regulator inspection dates. Check that regulators are in good condition. |
| Electrical equipment | Does it look in good condition? Is it all PAT Tested? If using 240 V, are there RCDs in place? |

| A3: Isokinetic sampling, wet chemistry and multiphase sampling |  |
| --- | --- |
| Pre-sampling |  |
| Type of test | e.g. Particulates. |
| Standard number | e.g. CEN, ISO etc. |
| Pitot leak check | Both Pitot openings should be checked for leaks before use. |
| Stack temperatures and pressures | These must be recorded on stack traverse sheet. |
| Angle of gas flow | Is a swirl test carried out? How is the angle of gas flow measured? The angle of gas flow should be less then 15° at each sample point. |
| Moisture | How was the pre-test moisture for pre-sample calculations obtained? Was it measured, estimated or based on historical data? |
| In-stack / out-stack | Out of stack testing is required if there are water droplets. |
| Sample duration | WIs, LCPs and BS EN 14181 specify minimum sample durations. If a batch process, how has the sample regime been decided? (e.g. peak emissions). |
| LOD | Check LOD is <10 % of ELV. |
| Sample blank | Observe / check a field blank has been carried out. |
| Equipment inspection |  |
| Equipment type | Usually manual sample train. Automatic kit used by some companies. |
| Sample train details | Is the glassware set up as per the standard? |
| Nozzles | Carry out visual check to ensure nozzles are sharp and round. |
| Probe length & materials | Usually titanium probe. Monel used for HF. High temperature work requires glass or stainless steel. Stainless steel not allowed for HCl sampling. |
| Pitots and calibration | Has the Pitot got a sharp edge? Is it labelled and calibrated? |
| Thermocouple calibrations | The following thermocouples are used: stack, probe, filter box, condenser, last impinger and gas meter Thermocouples must be calibrated. |

| A3: Isokinetic sampling, wet chemistry and multiphase sampling (continued) |  |
| --- | --- |
| Sample run |  |
| Pre-test leak check | Check for: kinked umbilical, umbilical caught under equipment, low flow or low vacuum. Are there any bubbles in the impingers? Expect to see meter box gas volume needle spin round rapidly then slow down or stop completely. Must be less than 2% of maximum flow rate (5% for dioxins). Ask the team leader what the leak check pass criteria is? Check / calculate if it is has passed. |
| Probe & filter temperature | Particulates: stack temperature or 160°C. Dioxins: filter temperature <125°C. HCl: > 150°C. Metals: 20°C above stack temperature or dew point (whichever is greatest). SO2: 120°C or 20°C above the (acid) dew point. |
| Impingers | Impingers should be kept cold with ice. |
| Condenser temperature (dioxins) | The dioxin trap must be less than 20°C. Thermocouple must be used to measure condenser outlet. |
| Mid-test leak check (dioxins only) | Must be carried out when changing port. Must be less than 5% of maximum flow rate. Must pass the leak check before sampling can be carried out. |
| Post-test leak check (dioxins only) | Must be carried out at the end of the sample run. Must be less than 5% of maximum flow rate. If leak check failed the sample is not valid. |
| Data recording | Must be recorded every 10 minutes. Check for dioxin condenser temperature cycles (i.e. the temperature should drop each time ice is added). |
| Reference conditions | Check with permit. If oxygen is measured, is the oxygen meter calibrated? |
| Post Test |  |
| Clean up, solutions and solvents | Check the required clean up solutions are used. Check they are applied in the required sequence. |
| Impinger weighing | Check weights should be used before weighing. |
| Sample labelling | Check each sample is uniquely labelled. |
| Sample storage | Do samples have to be kept in cool boxes or fridges? Do samples have to be kept in the dark? Is the sample storage temperature controlled? How long will the samples be stored before going to a laboratory? |

| A4: Instrumental methods |  |
| --- | --- |
| Pre-sampling |  |
| Type of test | E.g. CO by an infra red technique |
| Standard number | E.g. CEN, ISO etc. |
| Sample duration | WIs, LCPs and BS EN 14181 specify minimum sample durations. If a batch process, how has the sample regime been decided? (e.g. peak emissions). |
| Equipment inspection |  |
| Equipment type | This form is not suitable for FTIR. |
| MCERTS certification | Check if MCERTS certified. Check MCERTS certified range - 1.5 x daily ELV for WID, 2.5 x (daily) ELV for others. |
| Heated lines | Used for VOC, NOx and SO2. Check heaters are switched on and are at the required temperature. |
| Probe / filter type | Check appropriate materials are used. |
| Warm-up time | Have analysers been given enough time to warm up and stabilize? |
| Analyser calibration | Is the analyser within its calibration period? |
| Certified gases | Must be traceable to a gas certified by UKAS accreditation. Must have an uncertainty of <2%. |
| The test |  |
| Pre-test leak check | Must check full sample system |
| Pre-test zeros / span | Must check full sample system. Zero with N2 etc. Span at approximately 80% range (good practise to span at ELV) |
| Heated line temperature | Check it is at the required temperature. |
| Filter temperature | Must check full sample system. Zero with N2 etc. Span at approximately 80% range (good practise to span at ELV) |
| Post-test zeros / span drift | <2% acceptable. 2% to 5% needs adjusting. > 5% test invalid. |
| Data logging | Check the frequency of recorded readings and averaging periods. |
| Reference conditions | Most analysers measure dry gases. FIDs measure wet gases. |

| Site Specific Protocol (SSP) check sheet for stack emissions monitoring |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Contract/Job reference |  |  |  |  |  |  |  |  |  |
| Consultant |  |  |  |  |  |  |  |  |  |
| Operator |  |  |  |  |  |  |  |  |  |
| Installation name |  |  |  |  |  |  |  |  |  |
| Permit number |  |  |  |  |  |  |  |  |  |
| Report version number |  |  |  |  |  |  |  |  |  |
| Date of report issue |  |  |  |  |  |  |  |  |  |
| Requirement | Details/ items to be checked |  | Checked / x |  |  | Comments |  |  |  |
| Part 1 Contact details, monitoring dates and personnel | Operator name |  |  |  |  |  |  |  |  |
|  | Operator address and contact information |  |  |  |  |  |  |  |  |
|  | Permit number (including variations if applicable) |  |  |  |  |  |  |  |  |
|  | Installation name |  |  |  |  |  |  |  |  |
|  | Name of operator contact |  |  |  |  |  |  |  |  |
|  | Date and report number of previous monitoring campaign |  |  |  |  |  |  |  |  |
|  | Planned date of monitoring campaign |  |  |  |  |  |  |  |  |
|  | Name and address of monitoring organisation |  |  |  |  |  |  |  |  |
|  | Name, role during monitoring campaign, MCERTS registration number, certification level & technical endorsements held of all the persons to be involved in the monitoring campaign (certification expiry is to be included) |  |  |  |  |  |  |  |  |
| Part 2 Monitoring Objectives | The overall aim of the monitoring campaign |  |  |  |  |  |  |  |  |
|  | The substances to be monitored at each emission point -Emission Limit Value (ELV) - Reference conditions at which the results are expressed - Details of monitoring method(s) to be used for each substance |  |  |  |  |  |  |  |  |
|  | The overall uncertainty of the method(s) |  |  |  |  |  |  |  |  |
| Process Conditions | The type of process |  |  |  |  |  |  |  |  |
|  | A description of the process |  |  |  |  |  |  |  |  |
|  | If batch process, whether the whole of the batch is to be sampled or the details of the part of the batch sampled |  |  |  |  |  |  |  |  |
|  | The fuel type & feedstock |  |  |  |  |  |  |  |  |
|  | The normal load, throughput or continuous rating of the plant |  |  |  |  |  |  |  |  |
|  | Any unusual occurrences that take place during the process |  |  |  |  |  |  |  |  |
|  | What type of abatement system is fitted (if applicable) |  |  |  |  |  |  |  |  |
|  | What type of CEM system is installed and details of the data information system (if applicable) |  |  |  |  |  |  |  |  |
|  | The process details that need to be collected over the monitoring period |  |  |  |  |  |  |  |  |

| Requirement | Details/ items to be checked | Checked / x | Comments |
| --- | --- | --- | --- |
| Sample location | Dimensions of the stack(s) and monitoring facilities |  |  |
|  | A description of the location of the sampling plane for each release point |  |  |
|  | For each sampling plane, a description of the type of sampling port (accessibility, size, number, location) |  |  |
|  | For each sampling plane, a summary of the number, arrangement and orientation of the sampling line(s), and the number of sampling points per line |  |  |
|  | Access to the stack |  |  |
|  | Adequate work area at the sampling locations |  |  |
|  | Availability of required utilities (electrical, lighting, water) |  |  |
|  | Expected velocity profile* |  |  |
|  | Expected temperature & moisture of the stack gas* |  |  |
|  | Homogeneity test* (where required) |  |  |
|  | Restriction on using equipment, e.g. intrinsically safe areas |  |  |
|  | Physical restriction to using required apparatus |  |  |
|  | Appropriate measurement equipment for the application |  |  |
|  | For each sampling location, a summary of compliance with BS EN 15259 / EA TGN M1 (e.g. flow criteria, homogeneity, access to the sample line(s) and sample point(s) |  |  |
| Details of Monitoring | Expected emission values |  |  |
|  | The equipment used for each substance monitored |  |  |
|  | The sampling duration and number of samples for each measurement, including blanks |  |  |
|  | for manual methods, the proposed sample flow rate, volume and minimum sampling times |  |  |
|  | For instrumental methods, the proposed span gas concentrations (span gas should be 50-90% of analyser range or around the ELV) |  |  |
|  | The measurement concentration range and lower detection limit |  |  |
|  | For manual methods requiring a separate chemical analysis stage, details of the analytical method & the laboratory carrying out the analysis |  |  |
|  | Any modifications to the technical procedure, with justification and any resulting changes to the uncertainties |  |  |
|  | An explanation why any substance(s) in the monitoring objectives will not be monitored |  |  |
|  | An explanation why any substance(s) will not be monitored in accordance with the monitoring method |  |  |

* Historical information from a previous measurement reports may be used.

| Additional Information: TE1 - Particulates TE2 – phase partitioned species (metals, dioxins, PCBs, PAH) TE3 – gases and vapours by manual techniques TE4 – gases and vapours by instrumental techniques Additional comments |  |
| --- | --- |
| Summary and outcome |  |
| SSP checked by: |  |
| Date: |  |

| Check sheet for stack emissions monitoring reports |  |  |  |  |
| --- | --- | --- | --- | --- |
| Contract/Job reference |  |  |  |  |
| Monitoring organisation / contractors |  |  |  |  |
| Operator |  |  |  |  |
| Installation name |  |  |  |  |
| Permit number |  |  |  |  |
| Report version number |  |  |  |  |
| Date of report issue |  |  |  |  |
| Requirement | Details to be checked |  | / x | Comments |
| All Pages (Header/Footer) | Report reference number |  |  |  |
|  | Permit number (including variation number, if applicable) |  |  |  |
|  | Operator & installation name |  |  |  |
|  | Year of monitoring visit |  |  |  |
|  | Sequential number of the visit in the year (if applicable) |  |  |  |
|  | Version number |  |  |  |
|  | Page number (page x of y) |  |  |  |
| Cover Sheet (Not including Header/Footer) | MCERTS and UKAS logos and registration number of monitoring organisation |  |  |  |
|  | Permit number |  |  |  |
|  | Operator name |  |  |  |
|  | Installation name |  |  |  |
|  | Dates of monitoring visits |  |  |  |
|  | Contract number of reference |  |  |  |
|  | Name and address of client organisation |  |  |  |
|  | Name and address of monitoring organisation |  |  |  |
|  | Date of report |  |  |  |
|  | Name, MCERTS registration number and function of the person approving the report (must be Level 2 certified and hold the technical endorsements (TE) relevant to the substances being measured) |  |  |  |
|  | Signature of person approving the report |  |  |  |
| Contents | Description of parts 1 and 2 of the report |  |  |  |
| Monitoring Objectives | Overall aim of the monitoring campaign |  |  |  |
|  | Substances to be monitored at each monitoring point |  |  |  |
|  | Any special requirements stated |  |  |  |

| Requirement | Details to be checked | / x | Comments |
| --- | --- | --- | --- |
| Monitoring Results (for each emission point) | Emission Point Reference / Stack Reference |  |  |
|  | Substance to be monitored as permit and SSP |  |  |
|  | Emission Limit Value (ELV) expressed in the terms & units defined in the permit (e.g. mg/m3) |  |  |
|  | Periodic monitoring results completed for each emission point in the same terms as the ELV (e.g. mg/m3) and transcribed correctly from calculation sheets |  |  |
|  | Uncertainty associated with the result at a 95% confidence level |  |  |
|  | Units for the ELV, the periodic monitoring result and the uncertainty |  |  |
|  | Reference conditions at which the results are expressed (e.g. 273K, 101.3kPa 3% 02, dry) |  |  |
|  | Date of monitoring / sampling completed and transcribed correctly from calculation sheets |  |  |
|  | Start & end times of monitoring completed and transcribed correctly from calculation sheets |  |  |
|  | Sampling durations as per site specific protocol |  |  |
|  | Name & reference number of monitoring method used for each determinand |  |  |
|  | Accreditation for use of method (MCERTS, UKAS), for each determinand has been listed (check UKAS website for a schedule of accreditation for the monitoring organisation to verify) |  |  |
|  | Process / operating status at the time of monitoring |  |  |
| Operating Information | Whether the process was continuous or batch |  |  |
|  | Whether the whole of the batch was sampled or the details of the part of the batch sampled (if applicable) |  |  |
|  | What fuel was used during monitoring (if applicable) |  |  |
|  | What feedstock was used during monitoring (if applicable) |  |  |
|  | The normal load, throughput or continuous rating of the plant (if applicable) |  |  |
|  | Details of abatement system and whether operating |  |  |
|  | The periodic monitoring results and the results obtained for the corresponding period by the operator's CEMS completed and transcribed correctly from calculation sheets |  |  |
| Monitoring Deviations | An explanation why any substance(s) in the monitoring objectives was not monitored |  |  |
|  | An explanation why any substance(s) in the monitoring objectives was not monitored |  |  |
|  | Any other issues relevant to the monitoring results |  |  |

| Requirement | Details to be checked |  / x | Comments |
| --- | --- | --- | --- |
| Part 2 Appendix 1 | Name, role during monitoring campaign, MCERTS registration number certification level and technical endorsements held of the persons who carried out the monitoring campaign (are the technical endorsements required held by the team leader?) |  |  |
|  | Substances monitored, the ref number of the standard method used, and the reference to the technical procedure used by the monitoring organisation |  |  |
|  | A reference to the equipment checklists used on the monitoring campaign |  |  |
| Further Appendices (for each emission point) | Diagrams showing the dimensions of the stack and the monitoring facilities |  |  |
|  | Flow criteria measurements such as temperature, pressure and gas velocity Velocity difference highest to lowest <3:1 (9:1 for method 5 pressure readings) |  |  |
|  | Gas homogeneity test results |  |  |
|  | Gas measurements – oxygen, carbon dioxide (if required) |  |  |
|  | Water vapour measurements |  |  |
|  | Sampling measurements such as stack gas temperature and velocity during sampling |  |  |
|  | Instrumental gas analyser site calibration measurements inc zero and span gas concentrations (span gas should be 50-90% of analyser range or around the ELV) (if range has changed part way through, has the span check been redone?) |  |  |
|  | Instrument ranges stated and appropriate |  |  |
|  | Instrument drift stated and correct (0-5%) |  |  |
|  | Instrument results corrected for drift (if 2-5%) |  |  |
|  | Instrumental gas analyser results |  |  |
|  | Laboratory analysis sheets including the name of the analytical lab, accreditation for methods used and date of analysis |  |  |
|  | Manual methods results calculations, including concentrations and mass emissions (if required) |  |  |
|  | Isokinetic sampling achieved (95-115%) |  |  |
|  | Uncertainty calculations for each reported result, including |  |  |
|  | A list of parameters that affect the measurement uncertainty |  |  |
|  | Uncertainty sources (in order of significance) |  |  |
|  | Combined uncertainty for sources |  |  |
|  | Expanded total uncertainty as a percentage of the total measured |  |  |
|  | Expanded total uncertainty in units of measurement |  |  |
|  | Expanded total uncertainty as a percentage of the ELV (expressed at a 95% confidence interval) |  |  |

| Additional Information: TE1 - Particulates TE2 – phase partitioned species (metals, dioxins, PCBs, PAH) TE3 – gases and vapours by manual techniques TE4 – gases and vapours by instrumental techniques Additional comments |  |
| --- | --- |
| Summary and outcome |  |
| Report checked by: |  |
| Date: |  |